## Slide 1: The U.S Department of Energy Laboratory Accreditation Program & The Performance Testing Laboratory
Shane Steidley
Radiological and Environmental Sciences Laboratory

## Slide 2: Radiological and Environmental Sciences Laboratory (RESL)
  DOE-STD-1111-2018
  DEPARTMENT OF ENERGY LABORATORY ACCREDITATION PROGRAM ADMINISTRATION
    Authorized by DOE to conduct performance testing for DOELAP
    In Accordance with ANSI/HPS N13.30-2011
  DOE-STD-1112-2019
  DEPARTMENT OF ENERGY LABORATORY ACCREDITATION PROGRAM FOR RADIOBIOASSAY
    Points back to ANSI/HPS N13.30-2011 for performance testing criteria

## Slide 3: TESTING LABORATORY
  Radiological and Environmental Sciences Laboratory
    ISO/IEC 17025 General Requirements for the Competence of Testing and Calibration Laboratories
    ISO/IEC 17043 assessment - General requirements for proficiency testing
    ISO/IEC 17034 assessment - General requirements for competence of reference material producers
    ANSI/NCSLI Z540.2-R2002, U.S. Guide to the Expression of Uncertainty in Measurements

## Slide 4: TESTING LABORATORY
  NIST Traceability
  Application
  Test Plan
  Sample Preparation
  Analytical Verification
  Performance Evaluation

## Slide 5: Radiological Traceability Program (RTP)
Direct Traceability to NIST
  RESL prepares PE Materials for confirmatory measurements by NIST
  RESL analyzes PE Materials prepared by NIST
  All Matrices are addressed

## Slide 6: RTP Results (All Matrices)

## Slide 7: Testing Laboratory Application
Notes: Mixed Analyte PE materials

## Slide 8: Testing Laboratory Application
Notes: Mixed Analyte PE materials

## Slide 9: Testing Laboratory Comments
Notes: Application- Comments for assessors, you see something not specifically called out in ANSI standard (Tc-99 exa:0
Method needs such as sequential, …. Strings to pull for thoughts of assessment

## Slide 10: Testing Laboratory
PT tests on non standard ANSI N13:30
  Sm-141, Ni-63, Pu-241, ..
List MDAs for the nuclide of interest in the comments
Test samples based on ~10-20x MDA

## Slide 11: Testing Laboratory
Amounts of activity used in a given test shall be greater than or equal to the MTL and within the range specified in Table 2 or Table 3.
Notes: PT provider ensures keep within these ranges, for most part.
Some ANSI-revision discussion on ranges;
Two levels for sure we end up drilling into the application for more information: MDA’s, & system descriptions

## Slide 12: Testing Laboratory
How are the MTL’s determined
Established on capabilities of competent analysts using well-developed procedures and instruments that should, when properly quality controlled, give reliable results.
A MTL was originally developed based on approximately 10 times the MDA.
MTLs should not be construed with an acceptable MDA.

## Slide 13: Direct Radiobioassay Test Phantoms
Testing laboratory shall maintain and provide test phantoms, including appropriate blanks.
The testing laboratory shall identify the test category for each phantom but shall not divulge the activity or the identity of the radionuclide(s)
Kinda hard to do when a category is the radionuclide
The ability of the testing laboratory to measure quantitatively the amount of radionuclide placed in each phantom shall be affirmed by establishing traceability to NIST as described in ANSI N42.22 (1995).

## Slide 14: Direct Radiobioassay Test Phantoms
Testing laboratory instructs service laboratory to determine the amount of radionuclide in a phantom in a minimum of five replicate counts using the analytical procedure and counting times normally used for analysis of the radionuclide under consideration and to remove and reposition the phantom after each measurement.
Any deviations from the routine measurement protocol shall be documented in the report to the testing laboratory.

## Slide 15: RESL Analytical Verification
Testing laboratory verifies the amount(s) of the radionuclide(s) in the indirect radiobioassay test samples and phantoms.
Notes: Talk about how RESL does this.

## Slide 16: RESL Analytical Verification

## Slide 17: Plutonium Standard Cross Checks

## Slide 18
New Lungs- Since 2021
  Am-241 (3), MFP (3), Th-Equilibrium (2), Pitchblende (2)
  Am-241 (5), Am-243 (2), EuAm (1), Np (1), MFP (1), Nat U (1), U/Cs,/Co (4)
  Recertifying Lungs (12)
BOMABS
  K-40 (1), Eu-152 (2), MFP (5), Cs-134/137/Co-60 (1)
  2x (MFP, Eu) , 2x( Y, Sb, Eu), 2x (Cs, Eu)
  Multiple in production for Hanford & Japan
  Solid tissue substitute material available
Available Isotopes for SSPEP
  MFP (Co-60, Mn-54, Cs-134/137), K-40, Eu-152/154/155, Sb-125
  MDA for all types with variable counting times
Thyroids
  I-131, I-125 SSPEP available in January
  Ba-133, Sb-125 Available
DOE Phantom Library New
Notes: Blind testing or in-house calibrations (Phantom library availability), and resources

## Slide 19: Testing Laboratory Performance Reports
Notes: Results for Laboratory Name & associated lab code
No care about Log# RT retestings , Technical Equivelence
Direct: Lung, BOMAB, Thyroid (count 5 times and move/repostion and report 5 values)
Standard 13.30 =: Sqrt(B^2+Sb^2) = RMSE

## Slide 20: Testing Laboratory Performance Reports
Notes: Mixed Analyte PE materials

## Slide 21: Testing Laboratory Performance Reports
Notes: Six samples – Required to report 5 samples
Known values don’t have to be the same, but they can be.
Potential swapped sample

## Slide 22: Testing Laboratory Performance Reports
Notes: Six samples – Required to report 5 samples
Known values don’t have to be the same, but they can be.
Potential swapped sample

## Slide 23: Testing Laboratory Performance Reports
Notes: Six samples – Required to report 5 samples
Known values don’t have to be the same, but they can be.
Potential swapped sample

## Slide 24: Interesting Facts
Notes: Laboratory is on the money.

## Slide 25: Interesting Facts
Same Laboratory as before. Why is there a 7% bias between systems?

## Slide 26: Interesting Facts
Notes: Lab needs a CAR for PT value, you should be looking at this corresponding CA. Up to the DOLEAP administrator if it’s a good CAR or not.

## Slide 27: Interesting Facts
Same lab as previous. Shows they can do well on SU so what about fecal?

## Slide 28: Interesting Facts
RMC II Known to be biased ~10-15% High

## Slide 29: Interesting Facts
Wrong conversions (DPM-pCi).

## Slide 30: Interesting Facts
Statistical zeros
Using Wrong efficiency file