## Slide 1: Dosimetry DOELAP Assessor Training Review of Findings
Idaho Falls, ID
September 2024

## Slide 2: Writing Tips
Each finding should stand on its own merits.
Put the major issue up front.
Do not worry about upsetting the participant. Don’t write the excuses for a finding.
Ensure the write-up is complete and that a person who is not involved with the assessment (i.e. STM, OSB) can understand the issue from the write-up.

## Slide 3: Assessment Reports
Assessment Findings, Categorization, and Corrective Actions
  Your Responsibility:
    Given a finding, categorize it as an observation, concern, or deficiency
    If a concern or deficiency, tie the finding to a requirement
    “Should”, “recommend”, “suggest” or other words that imply assessor’s opinions tend to be observations
    Determine if site corrective action plan/response is adequate to close the finding
  What would you look for during the following site assessment?

## Slide 4: Corrective Action Review
STM sends CAP to assessors
  Brief review
  Big picture – did they address the issue you identified or saw?
  If they carry out the CA, will it fix the problem?
  Not a rock-bringing exercise
  Follow up will be assessed by subsequent assessors (or monitoring visit)

## Slide 5: DOELAP Findings Database
Chart: Test Session 2024 , 24 , 25 , 26 Assessment Finding
| 3.2(c) | 4.2(a) | 4.2(b) | 4.2(c) | 4.2(d) | 4.2(e) | 4.2(g) | 4.2(h) | 4.3(a) | 4.3(b) | 4.3(c) | 4.4(b) | 4.5(a) | 4.6(a) | 4.6(b) | 4.6(c) | 4.6(d) | 4.6(e) | 4.6(f) | 4.6(g) | 4.6(i) | 4.7(a) | 4.7(b) | 4.8(b) | 4.8(c) | 4.8(e) | 4.8(g) | 4.9(a) | 4.9(b) | App. B Paragraph 2 | Appendix C | DOE-STD-1111-2018 4.4.2
 | 1 | 1 | 1 | 1 | 1 | 2 | 3 | 4 | 1 | 1 | 3 | 8 | 2 | 2 | 5 | 2 | 9 | 5 | 4 | 4 | 2 | 4 | 2 | 1 | 1 | 1 | 1 | 2 | 3 | 1 | 6 | 1

## Slide 6: Example Finding #1
All of the procedures provided to the assessment team were currently up-to-date and reviewed. However, several procedures were discovered to have been reissued after exceeding the three-year review cycle timeframe as stated within the requirements of the Quality Assurance Plan. This appears to be a recurring issue. A concern on the 2017 onsite assessment found that eight procedures were not updated within the previously stated two-year review cycle timeframe. This concern was considered closed in 2021 during the “Quality Review for Programs in Good Standing” with the procedures being revised and officially reissued. However, these previously implemented corrective actions do not seem to have been effective in resolving this issue and do not adhere to the criteria in 4.4(b).
Concern

## Slide 7: Example Finding #1
Proposed Corrective Action C1a (CA-099659.01.01)
Analytical Services and Instrumentation (AS&I) and Technical Services Department (TSD) Management have agreed to provide additional administrative support to assist with the performance and coordination of the procedure review and revision process. Assistance may include:
•Maintaining a procedure tracking index/log and procedure revision schedule.
•Providing review and revision notifications to AS&I personnel and management.
•Ensuring the final procedures are properly formatted and updated with information provided by technical staff.
•Obtaining the appropriate procedure signatures.
•Electronically distributing, posting, and controlling the documents for user access.
Proposed Corrective Action C1b (CA-099659.01.02)
Procedure review dates may be adjusted to reduce the need to review and reissue multiple procedures at the same time. However, the review reissue dates shall not be extended past the review cycle timeframe as specified in the Radioanalytical Laboratories Quality Assurance Plan (i.e., the procedure reissue date may not be extended past 3 years from the current issue date).
Proposed Completion Date: 7/1/2024

## Slide 8: Example Finding #2
Between April 1, 2020, through March 31, 2021, there were four incidents involving at least 13 QC samples and 45 worker samples that were reported as failed analyses due to problems during the filtering process. However, no corrective action request and report (CARR) or an internal nonconformance report were opened to investigate and prevent a recurrence. In June 2021, another filtering problem resulting in sample failures was finally investigated and corrective actions implemented. Section E in GEL's Quality Assurance Plan (GL-QS-B-001, rev 36) requires the completion of an internal nonconformance report when laboratory control samples do not meet acceptance criteria specified in the applicable analytical SOP. DOE-STD-1112-2019 4.6 (b).
Concern

## Slide 9: Example Finding #2
Response: This concern is addressing data sent over a full year period and indicates that a trend was not addressed by a CARR. Failed analyses are addressed in the quarterly and annual reports provided to DOE clients. The review of the quarterly and annual reports would be the best time to identify repeat issues that should be addressed with CARRs.
Proposed Actions #1: A Quality review of the quarterly/annual reports will be added so that trends could be identified at that time since individual emails of failed analyses events do not indicate trends.
Proposed Actions #2: When a failed analysis occurs, it will be entered into the laboratory’s NCR system for tracking and investigation. The LIMS system can be queried to determine repeat issues of failure. In the instances of repetitive causes of failure, a Corrective Action would be initiated. In addition, a Quality review of the quarterly and annual reports will be added to each workorder in LIMS so that the effectiveness of the Corrective Actions can be evaluated at that time.

## Slide 10: Example Finding #3
XXX utilizes a system named Document Organization and Control System (DOCS) for procedures used during the course of their work. While DOCS is adequate for tracking purposes, and captures revisions well, it is slightly confusing to determine procedure revision dates for procedures when only reviews take place, and no actual changes are made to previous versions. Other issues include that in this system the only person able to make changes is the owner of the document, effectively locking out other procedure users needing to make changes, and personnel named on the DOCS automatically generated cover sheet may be out of date. However, since this is a site-wide system, making global changes to the system application is out of the control of Internal Dosimetry staff.
Observation

## Slide 11: Example Finding #4
Section 3.2 (c), MDA units for urine matrix on DOELAP application and SOW with GEL are not consistent: DOELAP application lists MDA units as pCi/Liter while the GEL SOW lists MDA units as dpm/Sample.
Observation

## Slide 12: Example Finding #5
Acceptance criteria for GEL QC data is not clearly defined in the SOW with GEL; furthermore, the SOW infers that failures for QC results do not necessitate corrective actions. No requirements are included in the SOW for re-evaluation of all measurements since last acceptable QC. (QI7, QI8)
Concern
4.6(b)
Cause Rationale: General Environmental Laboratories (GEL) performs acceptance of Quality Control (QC) data in accordance with their procedures however that process was not clearly defined in the scope of work.
  Revise GEL Scope of Work (SOW) Quality Control Samples Section to clearly define acceptance criteria for QC data.
  9/28/2023

## Slide 13: Example Finding #6
The assignment of the quality lead role (however named) and the responsibilities for this role are not clearly defined in documentation. [4.2(c)] This role could also be considered ‘key personnel'''' and would therefore not meet the requirements of 4.2(d).
Concern
Response: The overall quality of the Radiobioassay program is included in the responsibilities of the Dosimetry Program Manager (Technical Lead as the position is identified in the Radiobioassay Standard DOE-STD-1112-2019, 4.2(b) of appendix B of XXX-4007 Internal Dosimetry Program Quality Assurance Plan.
CAP: Revise appendix B of XXX-4007 Internal Dosimetry Program Quality Assurance Plan to separate out the responsibilities of the quality assurance lead. Note that the function may be performed by the technical lead or delegated by the Dosimetry Program Manager to another professional level Health Physicist or Radiological Engineer.

## Slide 14: Example Finding #7
Procedures for acceptance testing OSL dosimeters have not been finalized and are still in draft form. 4.5(a), 4.7.1(b), 4.7.1(e).
Concern
Issue procedures for acceptance testing OSL dosimeters. Completion Date: 6/10/2024
Finalize and publish the drafted procedure for acceptance testing OSL dosimeters. Including the routing through the company’s formal review, verification and validation, approval, and publication processes. Completion Date: 6/10/2024
Internal QA “policy” on minimal CAs

## Slide 15: Example Finding #7
Supplemental Information
The Landauer Dosimetry System and the associated dosimeters were purchased in late Fiscal Year 2018. At that time, the entire system including the dosimeters were accepted. In 2019 and continuing into 2020, very preliminary investigations into the working of the Landauer Dosimetry System were performed. Near the end of 2020, it was realized that the dosimeters provided to the NNSS were not the appropriate type and Landauer was contacted. Over the next several months, the appropriate dosimeters were provided by Landauer and accepted by MSTS. These dosimeters constitute the current population of OSL dosimeters ready for use with the Landauer Dosimetry System and are adequate for several years of operations.
Realizing the current supply of OSL dosimeters is sufficient negated the urgency to develop a formal dosimeter acceptance procedure. As testing and development of the Landauer Dosimetry System progressed a draft procedure was developed and tested by NNSS to ensure the acceptance testing procedure is available once the system is accredited and before any additional dosimeters are procured. The draft procedure is going through the rigorous review process required by CD-3500.000, Directives Management System. The review process requires a formal review by specific personnel (including the Dosimetry Supervisor, Dosimetry Manager, and others), performance of validation and verification by qualified Dosimetry Technologists, and resolution of any resulting comments. The CAP specifies the procedure must be issued with a completion date of June 10, 2024.

## Slide 16: Example Finding #8
The analytical results reports have the initials of the reviewing technical staff, but the identity of the reviewers cannot be determined, and the document reports are not date-stamped to determine who and when the reviews were completed. The software used in the In-vivo calculations also needs to be listed on the report for reproducibility. Lastly, the In-vivo documents and instrument labels note just 'the year' of review/calibration not the day and month the review was completed.
Observation

## Slide 17: Example Finding #9
The procedure 12-DS-1341, Rev. 11 was revised dated 03/15/2021 to rename the procedure as "Internal/External Dosimetry Program Accreditation Application and lntercomparison" as mentioned in the XXX Response for corrective action. However, it does NOT address the main concern to include: "the actions for entering the DOELAP findings into a commitment tracking system (CTS)" to resolve the intent to track and close the concern.
Note that this was Concern 1 in the previous assessment so elevated to a deficiency.
Standard citation? Original Concern didn’t have citation either.
  Maybe 4.1(b) ?

## Slide 18: Example Finding #10
ZZZ performs annual proficiency evaluations that, by procedure, can include on the job training, testing, or observation of operations. As such, the evaluations do not specifically require annual observations of operations as required by DOE-STD-1112-2019.
CONCERN, 1112-2019 4.3(c) The technical lead shall initially and at least annually evaluate and document the proficiency of each staff member authorized to perform radiobioassay functions. This proficiency assessment shall include an observation of performance.
1095-20XX 4.4(d) The Technical Lead shall initially and at least annually evaluate and document the proficiency of each staff member authorized to perform dosimetry related functions. When appropriate, this proficiency assessment should include an observation of performance.

## Slide 19: Example Finding #10
The procedure for 237Np analysis in urine and feces is not sufficient for separating interferences and resolving a mixture of radionuclides. The Project Lead estimated that about 25% of the 237Np urinalyses required additional cleanup after counting. Though GEL passed the DOELAP performance test for 237Np in urine, the procedure should still be evaluated for improvement. DOE-STD-1112-2019, 4.6 (c).
Concern
DOE-STD-1112-2019, 4.6 (c) The laboratory shall use appropriate techniques to ensure the proper identification and qualification of specific radionuclides, including separating interferences and resolving a mixture of radionuclides.

## Slide 20: Example Finding #11
GL-DC-E-001; Document Control procedure section 6.1.5 stated that each level of review is dated and documented by signature on the cover of the published procedure. Contrary to GL-DC-E-001, the procedures available to assessors and general staff do not show the last review date and signature on the cover pages. Assessors were informed this is due to security concerns associated with distributing a digital copy of signatures. Review dates are available on the table of contents associated with the different procedure types.
Observation

## Slide 21: Example Finding #12
XXX documents that form the quality assurance program are not reviewed for accuracy in accordance with documented internal review frequencies. XXX- 55719, Rev. 1, XXX Internal Dosimetry Program Manual, was issued in January 2017. The document has not been included into a 3-year review cycle for technical procedures or 5-year cycle for administrative procedures. Because it has not been recently reviewed it does not reference the current version of DOE-STD-1112. DOE-STD-1112-2019, 4.4(b)
Concern

## Slide 22: Example Finding #13
Calibrations for direct radiobioassay instrumentation are not conducted or reviewed in accordance with established procedures. Proof of review and completion for the Quality Assurance File, Technical Evaluation, Instrumentation Event Log, and Energy/FWHM Calibration Results were not reviewed and completed per referenced procedures and processes.
Concern
DOE-STD-1112 4.8(c)

## Slide 23: Example Finding #14
Criteria 4.9(b) of the DOE STD-1112-2019 addresses that the indirect radiobioassay measurement report developed for permanent record shall include values of the decision levels in the same units as the results. The indirect gamma spectroscopy bioassay measurement report does not include values of the decision levels or notification levels. Although the gamma measurement report provides the MDA values, the official permanent record also needs to include decision levels values.
Concern

## Slide 24: Example Finding #14
Proposed Corrective Action C4a (CA-099659.04.01)
Lack of decision level (Lc) is a recognized and longstanding technology shortfall (as defined in DOE-STD-1111-2018 ) of the Mirion APEX Gamma software. The shortfall has been documented by the indirect radiobioassay program in multiple instances, the most recent being memo 2022-003.
The indirect gamma spectroscopy bioassay measurement report does include the customer specified notification level in the same units as the results. The information in the Laboratory Information Management System (LIMS) is part of the official permanent record per the Radioanalytical Laboratories Quality Assurance Plan Rev 1.3, Section 4.5.2 c. The LIMS Notification Level table lists the notification level as MDA and the units as uCi. The Results Summary field in LIMS displays this information for each sample/analyte combination for all gamma analytes identified in a given sample.
To address the Lc technology shortfall, the Lc for gamma spectroscopy analyses shall be estimated as MDA/2 for all reported radionuclides. The basis for this estimate is derived from the equations found in Curie (1968) and is documented in the technical justification document 2023-114 Justification for Estimating Lc from the Reported MDA for Gamma Spectrometry Analyses. The estimate of the Lc shall be hand-entered into the LIMS using the same process as that described for hand entering the MDA in Section 6.5 of Preparation of Bioassay Urine and Fecal Samples for Analyses by Gamma Spectroscopy.

## Slide 25: Example Finding #14
Proposed Corrective Action C4b (CA-099659.04.02)
Preparation of Bioassay Urine and Fecal Samples for Analyses by Gamma Spectroscopy shall be updated to include Lc calculations and LIMS data entry instructions for the Lc.
Proposed Completion Date: 2/1/2024

## Slide 26: Example Finding #15
As several staff members in critical roles are nearing retirement, XXX has proactively hired additional staff for succession planning. These additional staff members are currently being mentored and/or cross-trained by experienced personnel to ensure that there is adequate knowledge transfer, and ultimately results in no disruption in service or loss of quality when these experienced employees leave XXX.
Observation – Noteworthy practice

## Slide 27: Example Finding #16
C1: Section 4.6 (b): In vitro quarterly reports are not completed in a timely manner. Per procedure DP-6.04, Coordination of the Excreta Quality Control Oversight Program, quarterly reports shall be completed within three months after in vitro data is received from GEL. The April 1st through June 30th, 2021 quarterly report was finalized on May 6th, 2022 and the GEL data was received August 30th, 2021. Weekly reviews of radiobioassay results from GEL do not include a quality assurance review by the Quality Assurance Engineer
Concern

## Slide 28: Example Finding #16
Corrective Action Plan for Concern #1
– ACTION: (ITS A-02920-001) Institute formal action tracking of due dates for quarterly quality control data receipt from GEL (45 business days following the end of each quarter, per GEL contract SOW)
– ACTION: (ITS A-02920-002) Institute formal action tracking with due dates for completion of quarterly Quality Control Review Report (3-months following receipt of data from GEL, per XXX contract SOW).
– ACTION: (ITS E-02920-003) Institute reviews by the Quality Assurance Engineer of the weekly excreta review report that includes the radiobioassay data received from GEL.