## Slide 1: Dosimetry DOELAP Assessor Training Review of Findings
Idaho Falls, ID
September 2026

## Slide 2: Writing Tips
Each finding should stand on its own merits.
Put the major issue up front.
Do not worry about upsetting the participant. Don’t write the excuses for a finding.
Ensure the write-up is complete and that a person who is not involved with the assessment (i.e. STM) can understand the issue from the write-up.

## Slide 3: Report Assessment Workshop
Assessment Findings, Categorization, and Corrective Actions
  Your Responsibility:
    Given a finding, categorize it as an observation, concern, or deficiency
    If a concern or deficiency, tie the finding to a requirement
    Determine if site corrective action plan/response is adequate to close the finding
  What would you look for during the following site assessment?
  1095-2025 4.5: The corrective action plan shall include the actions to be taken by the program to address the concerns and deficiencies, as well as the dates of completion of the actions.

## Slide 4: Corrective Action Review
Corrective action plans submitted to the STM, through the DOE Field Office, within 45 days of closeout (DOE-STD-1111)
Very “optimistic”
STM sends CAP to assessors
  Brief review
  Big picture – did they address the issue you identified or saw?
  If they carry out the CA, will it fix the problem?
  Not a rock-bringing exercise
  Follow up will be assessed by subsequent assessors (or monitoring visit)

## Slide 5: Finding example #1
The training documentation is in paper form and not maintained electronically for dosimetry-related tasks. There is no process in place to remind personnel that training is due or alert when training is out of date. Some training documents referenced previous-generation dosimetry systems. The training program needs to be updated, maintained in a more rigorous fashion, and a system implemented to ensure personnel are current in the appropriate training areas and requirements. (TQ.1, TQ.2, TQ.3).
Concern
References?
4.4(a), 4.4(b), 4.4(c)

## Slide 6: Finding example #1 proposed corrective actions
Revise EDL-02, Dosimetry Training, to reflect an updated internal training program and training records management system. Include, as appropriate, updates or clarification for the following training program elements: (1) Initial written evaluations for each qualification type and level. (2) Examples of tasks suitable for annual performance observation requirements. (3) Description of required reading program for initial training, annual retraining, and upon procedure revision.
Include a rigorous training records management system to ensure signed paper files are digitized and originals are appropriately protected in accordance with XXX Institutional Procedure PRO-0207, Managing Records. The second part of this corrective action addresses Observation #3 as identified in the assessment report. Although responses to observations are not required by DOELAP, XXX agrees with the observation and will take the corrective action identified.
Develop an electronic training management system to track dosimetry program training requirements, completions, and due dates. Ensure new requirements for training records management described in the revision to EDL-02, Dosimetry Training, are implemented.

## Slide 7: Finding example #2
Annual observation of performance of personnel is not being performed. XXX training requalification periods have been set to a two year interval to match plant standards. No observation of performance is being performed in the interim years. (TQ4)
Concern
References?
1095-2025 4.4(d) The Technical Lead shall initially and at least annually evaluate and document the proficiency of each staff member authorized to perform dosimetry-related functions. When appropriate, this proficiency assessment should include an observation of performance.

## Slide 8: Finding example #2 proposed corrective actions
All Dosimetry Qualification cards have been updated to reflect an annual requalification period in order to document personnel performance sufficiently. This action is complete.

## Slide 9: Finding example #3
The Blind Audit program is established for both whole-body and extremity dosimeters. The DOELAP expectation for a blind audit program under the new standard was recently communicated to provide clarification for requirements TD20, OD20, and TE18 and stated “If accredited in IIA, program must include two x-ray beams from 20 to 70 keV and at least two from 70 to 500 keV. 137Cs and/or 60Co shall always be included. If accredited in IIIA, then include 90Sr/90Y and 85Kr. If accredited in VA, programs shall include both bare and moderated 252Cf.” The XXX blind audit program currently does not include this specificity (TD20, OD20, and TE18).
Concern
References?
1095-2025 4.8.2(f)

## Slide 10: Finding example #3 proposed corrective actions
XXX will update the blind QC procedure to include the required source specificity. XXX will update any affected contracts to ensure that blind QC irradiations are performed for all required sources. Estimated completion date: November 30, 20XX

## Slide 11: Finding example #4
The XXX internal dosimetry procedure RC-DOS-3 Rev 18 establishes and provides guidance for the implementation of a Blind Performance Testing Program (BPTP). However, the BPTP is technically lacking in several areas. In general, the blind testing samples that are submitted for analyses, over time, shall include samples that demonstrate accuracy, false positive, false negative and sensitivity evaluations. Furthermore, the blind testing samples shall contain chemical, matrix and radionuclide interferences common to the program's routine samples. While XXX and the YYY radiobioassay programs have similar analytical needs and the YYY provides XXX their data which is reviewed by both programs, all the areas of testing are not being met. For example, sensitivity testing using activity concentrations slightly above minimum detectable amounts and false negative testing using activity concentrations that are present in the sample that are slightly above detection limits are not being performed.
Using the stated RMSE acceptance criteria, for acceptance testing, is not technically acceptable as RMSE acceptance criteria should only be used for accuracy testing, on multiple samples, and does not apply to individual results. A BPTP would have the capability to analyze the testing results on an individual basis so that if a blind audit sample is submitted with a group of worker samples and the blind audit sample has unacceptable results then only that group of samples would be reevaluated instead of all the measurements that were made since the last acceptable results. Reevaluation of all XXX measurements after failures resulting from nonacceptable quarterly RMSE acceptance criteria is not being performed at XXX.
The YYY data reports do contain some 2 sigma evaluations, but they are being interpreted incorrectly and there are examples that if the test being performed failed due to being outside of the 2 sigma evaluation, that they were written off as being acceptable because the results fell within 3 sigma. This practice is unacceptable as a healthy BPTP would simply note the failure, investigate the cause of failure, and concentrate supplemental monitoring in the area of the failure and if the subsequent results continue to indicate problems, then more rigorous actions would be taken. There is no evidence, in the 1st and 2nd 2022 quarterly reports that any subsequent actions of further testing were taken for failures originating within YYY samples CFA22001, CFA22003, CFA21086 and CFA21084.
Data the XXX is using for their BPTP from the YYY quarterly reports for the 1st and 2nd quarters of 2022 indicate issues with 7 out of 18 radionuclides evaluated by the YYY in both urine and fecal samples for the first quarter, and 3 issues out of 8 radionuclides evaluated by the YYY in both the urine and fecal samples from the second quarterly report. These indications show that there are problems with almost 50% of the analyses of samples report to the YYY by ZZZ. Though they may not be applicable to the XXX, they may demonstrate that it would be beneficial to concentrate monitoring efforts on a more frequent schedule until the analytical issues are resolved (4.6(b), 4.6(c), 4.6(d), 4.6(e)).

## Slide 12: Finding example #5
Reviewing the previous assessment:
Concern 2: A separate Panasonic reader is located at an off-site location as a means to ensure processing capabilities. However, this off-site reader has not been calibrated. The backup processing procedure does not reflect the current configuration. (E.5)
Status: The Panasonic TLD reader was calibrated and made available for use of a backup system; however, calibration was not maintained during the accreditation period. The system failed calibration in May 2022. The unit was correctly taken out of service; however, it was not recalibrated and placed back into service until August 19, 2022.
What to call this?
This finding remains open and will be elevated to Deficiency 3-1.
References?
1095-2025 4.2.1(b)

## Slide 13: Finding example #5 proposed corrective actions
XXX will maintain its back up dosimetry reader in calibration and in service. Calibration requirements for the Panasonic TLD will be institutionalized into XXX’s XX-109, Instrument Data and Recall Tracking System. The readers will have a 6-month calibration recall for the photon counter and a 12-month calibration recall for the frequency counter. If a reader were to become unavailable, the reader will be returned to service within a 14-day period by either XXX’s service provider or by XXX internal resources. XXX will communicate with the DOELAP Senior Technical Manager if the reader cannot be restored in 14-days.

## Slide 14: Finding example #6
Performance data for the Harshaw HBG dosimeter in Test Category IIIA (Beta - General) indicates a clear, highly reproducible (S = 0.022) negative bias (B = -0.162). It is recommended that the source of this bias be investigated, and if identified, corrected.
Observation
References?
No response needed

## Slide 15: Finding example #7
Although the PMDP passed all Performance Testing categories, they were prompt in initiating evaluation and corrective action to address an overall negative bias (<30%) in Test Category IVA-1 (photon + Kr-85). This is a noteworthy practice.
Observation
Please continue to highlight noteworthy practices, but preferred to write them in the Comments section.

## Slide 16: Finding example #8
XXX returns control dosimeters with the field dosimeters for making the background subtraction. A table for control limits was created to evaluate this subtraction, with these limits based on a function related to a daily background rate. However, there was no evidence that supported the accuracy of this background level. XXX should confirm the daily background rate used. This is a particularly important step in identifying potentially anomalous readings.
Concern
References?
1095-2025 4.8.1(c)

## Slide 17: Finding example #8 proposed corrective actions
Corrective Action Plan for Concern #1
C.1.1-ACTION: Conduct a background evaluation/study from historical dosimetry data and formally document the results in a revision to the External Dosimetry-Technical Basis Document (ED-TBD), due by October 19, 20XX. (Attach the approved revision as objective quality evidence.)
C.1.2-ACTION: Draft a procedure revision to DP-3.01, Review of Landauer Dose Results, Exhibit 5, Acceptable Range for Landauer Control Doses consistent with the revised ED-TBD, due by November 2, 20XX. (Attach draft procedure change as objective quality evidence.)
C.1.3-ACTION: Finalize and publish procedure revision to DP-3.01, Review of Landauer Dose Results, Exhibit 5, Acceptable Range for Landauer Control Doses, due by November 28, 20XX. (Attach final procedure as objective quality evidence.)

## Slide 18: Finding example #9
Although whole-body dosimeters are inspected to ensure that all necessary components are in place before being placed into service, procedures do not require inspection of Saturn ring dosimeters. 4.8.1(e).
Observation

## Slide 19: Finding example #10
Work processes for accreditation activities must include stated acceptance criteria and inspection points. The review and approval of the dose data from Landauer is not formally documented, and the procedure does not include stated acceptance criteria for control dosimeter data or personal doses. There are unclear acceptance criteria and no documented reviews of background dosimeters results. The inspection of the lnlight dosimeters upon receipt from Landauer does not include a verification that CR39 is present in the dosimeter case, i.e., that the dosimeter conforms to the ordered product. [4.5(a),4.7(b), 4.7.1(e)]
Concern
1095-2025 4.6(a), 4.8(b), 4.8.1(e)

## Slide 20: Finding example #10 proposed corrective actions
Corrective Action Plan:
(1) Revise drafted new Procedure PROC-RP-4587, Routine Dosimeter Exchange and Manual assignment of Dosimeters, to include of inspections of the dosimeter shipments, dosimeter parts/components and verification that CR39 is present in the dosimeter cases.
Due Date: 09/29/2022
(2) Draft new procedure to document review of dose data from Landauer prior to being loading into the dosimetry records system. Procedure to include acceptance criteria for control, personnel, and background dosimeters.
Due Date: 11/30/2022