## Slide 1
(1) A formal review of the QA program, including elements specified in DOE-STD-1095-2018, must be conducted at the midpoint of the DOELAP assessment cycle. This formal review, referred to as a Management Assessment, is required in PROCEDURE-QA-1, and is scheduled for the summer of 2022. However, as of this DOELAP onsite assessment, it has not been completed.

## Slide 2
(1) A formal review of the QA program, including elements specified in DOE-STD-1095-2018, must be conducted at the midpoint of the DOELAP assessment cycle. This formal review, referred to as a Management Assessment, is required in PROCEDURE-QA-1, and is scheduled for the summer of 2022. However, as of this DOELAP onsite assessment, it has not been completed. [4.2(g)] CONCERN

## Slide 3: The performance tests in the accreditation categories were passed but all the categories had a negative bias for last two test sessions for Whole-body and Saturn rings. The standard requires that the applicant shall review the performance testing data for potential improvement in the dosimetry measurement system.

## Slide 4: The performance tests in the accreditation categories were passed but all the categories had a negative bias for last two test sessions for Whole-body and Saturn rings. The standard requires that the applicant shall review the performance testing data for potential improvement in the dosimetry measurement system. [3.2(e)] CONCERN

## Slide 5: No backup performance testing has been performed over the past three years to demonstrate or verify the effectiveness of backup service provider capabilities to ensure dosimetry results remain accurate, repeatable, and verifiable. Backup service provider performance testing is expected to be exercised at a frequency identified in the QAP to ensure dosimetry results are accurate, repeatable, and verifiable; however, this frequency is not defined within the current QAP.

## Slide 6: No backup performance testing has been performed over the past three years to demonstrate or verify the effectiveness of backup service provider capabilities to ensure dosimetry results remain accurate, repeatable, and verifiable. Backup service provider performance testing is expected to be exercised at a frequency identified in the QAP to ensure dosimetry results are accurate, repeatable, and verifiable; however, this frequency is not defined within the current QAP. OBSERVATION Concern? 4.2.1(c), 4.2.1(d)

## Slide 7
(3) The Dosimetry Vendor states that other worldwide locations provide for continuity of operations. The program does not have a documented plan but should consider how dosimetry operations may be affected by the unexpected loss of key personnel.

## Slide 8
(3) The Dosimetry Vendor states that other worldwide locations provide for continuity of operations. The program does not have a documented plan but should consider how dosimetry operations may be affected by the unexpected loss of key personnel. OBSERVATION

## Slide 9
(4) The QAP, Technical Basis, and the current contract specification, in several places, refer to the ‘DOELAP accreditation’ of the Dosimetry Vendor. These statements should be corrected – only DOE programs are accredited by DOELAP.

## Slide 10
(4) The QAP, Technical Basis, and the current contract specification, in several places, refer to the ‘DOELAP accreditation’ of the Dosimetry Vendor. These statements should be corrected – only DOE programs are accredited by DOELAP. OBSERVATION

## Slide 11
(5) The DOE contract for the management of cleanup activities on the DOE XYZ Site was awarded to ACME just two days prior to the start of this onsite assessment. External dosimetry program procedures and desktop instructions have been reissued as-is with a cover document that addresses the new organization and new position titles. The program has been granted a 1-year period to make modifications to the procedures to formally incorporate the new organizational structure.

## Slide 12
(5) The DOE contract for the management of cleanup activities on the DOE XYZ Reservation was awarded to ACME just two days prior to the start of this onsite assessment. External dosimetry program procedures and desktop instructions have been reissued as-is with a cover document that addresses the new organization and new position titles. The program has been granted a 1-year period to make modifications to the procedures to formally incorporate the new organizational structure. OBSERVATION

## Slide 13
(6) Work processes for accredited activities must include stated acceptance criteria and inspection points. The review and approval of dose data from the vendor is not formally documented, and the procedure does not include stated acceptance criteria for control dosimeter data or personnel doses. There are unclear acceptance criteria and no documented review of background dosimeter results. The inspection of dosimeters upon receipt from the vendor does not include a verification that CR39 is present in the dosimeter case, i.e., that the dosimeter conforms to the ordered product.

## Slide 14
(6) Work processes for accredited activities must include stated acceptance criteria and inspection points. The review and approval of dose data from the vendor is not formally documented, and the procedure does not include stated acceptance criteria for control dosimeter data or personnel doses. There are unclear acceptance criteria and no documented review of background dosimeter results. The inspection of dosimeters upon receipt from the vendor does not include a verification that CR39 is present in the dosimeter case, i.e., that the dosimeter conforms to the ordered product.
OBSERVATION
Concern? [4.6(a), 4.8.1(b), 4.8.1(e)]

## Slide 15
(7) Blind testing program irradiation schedule includes Category IIA for which XYZ is accredited for the whole-body dosimeter with CR-39 and the extremity ring dosimeter. However only Cs-137 and Am-241 are tested for this category. DOELAP expectations for Category IIA are that x-ray sources used in blind testing will include at least two beam codes with energies between 20 keV and 70 keV, and two between 70 keV and 300 keV, in addition to either Cs-137 or Co-60.

## Slide 16
(7) Blind testing program irradiation schedule includes Category IIA for which XYZ is accredited for the whole-body dosimeter with CR-39 and the extremity ring dosimeter. However only Cs-137 and Am-241 are tested for this category. DOELAP expectations for Category IIA are that x-ray sources used in blind testing will include at least two beam codes with energies between 20 keV and 70 keV, and two between 70 keV and 300 keV, in addition to either Cs-137 or Co-60. [4.8.2(e)] CONCERN

## Slide 17
(8) Dosimeter system characteristics are documented via vendor provided technical reports. However, some of these system characteristics are not documented or referenced in the XYZ External Dosimetry Technical Basis Document.

## Slide 18
(8) Dosimeter system characteristics are documented via vendor provided technical reports. However, some of these system characteristics are not documented or referenced in the XYZ External Dosimetry Technical Basis Document. [4.8.1(c), 4.8.1(d), 4.8.1(g)] OBSERVATION

## Slide 19
(9) In preparation for the announced retirements of one External Dosimetry Technical Specialist and the program Director, successor personnel are in place to work alongside the experienced personnel to ensure a successful transition. This is considered a noteworthy practice.

## Slide 20
(9) In preparation for the announced retirements of one External Dosimetry Technical Specialist and the program Director, successor personnel are in place to work alongside the experienced personnel to ensure a successful transition. This is considered a noteworthy practice. OBSERVATION

## Slide 21
(10) There is no documented training program for the External Dosimetry Technical Specialists and the EDP Manager. They can perform accredited activity in various facets of operation, including review of dosimeter processing results, operation of the readers, and assigning dose to an individual.

## Slide 22
(10) There is no documented training program for the External Dosimetry Technical Specialists and the EDP Manager. They can perform accredited activity in various facets of operation, including review of dosimeter processing results, operation of the readers, and assigning dose to an individual. (4.4(a) through 4.4(c)). CONCERN

## Slide 23
(11) The External Dosimetry Management System is the main database repository system for the XYZ dose of record. The report and forms used to access the data tables are planned to be rewritten in Cold Fusion, as their current versions (Oracle Forms) is nearing end-of-life. This will be a long-term, large-scale endeavor.

## Slide 24
(11) The External Dosimetry Management System is the main database repository system for the XYZ dose of record. The report and forms used to access the data tables are planned to be rewritten in Cold Fusion, as their current versions (Oracle Forms) is nearing end-of-life. This will be a long-term, large-scale endeavor. OBSERVATION
Future Technical Equivalency?

## Slide 25
(12) The RCF for DXT calibration performed on May 6, 2022 for an ECC read was outside of the posted normal range for reader 7 but was accepted for use by the ED Technician without reporting to the EDP Manager. Procedure EDP-607, Daily Reader Preparation prescribes reporting these results to the EDP Manager for review. (The EDP Manager may approve any RCF that is within 10% of the mean of the current RCF range.

## Slide 26
(12) The RCF for DXT calibration performed on May 6, 2022 for an ECC read was outside of the posted normal range for reader 7 but was accepted for use by the ED Technician without reporting to the EDP Manager. Procedure EDP-607, Daily Reader Preparation prescribes reporting these results to the EDP Manager for review. (The EDP Manager may approve any RCF that is within 10% of the mean of the current RCF range. [4.6(a) and 4.8.2(b)] CONCERN

## Slide 27
(13) The backup processing services procedure was tested in late 2020 with a small population of HBG dosimeters. More frequent testing for all dosimeter types including larger numbers of dosimeters is recommended to ensure all aspects of the process are adequate in the event that both readers are not operational.

## Slide 28
(13) The backup processing services procedure was tested in late 2020 with a small population of HBG dosimeters. More frequent testing for all dosimeter types including larger numbers of dosimeters is recommended to ensure all aspects of the process are adequate in the event that both readers are not operational. [4.8(c)] OBSERVATION

## Slide 29
(14) DOE-STD-1095 states that Quality Control and unirradiated dosimeters shall be used routinely to identify reader processing problems. Unirradiated and quality control dosimetry frequency shall be determined based upon the total number of dosimeters processed, equipment stability, type of quality control checks, or other suitable method. The practice used by XYZ (XYZ procedure QA-1-009 step 7.2.7) is that only irradiated Quality Control cards (QC) are interspersed in dosimeter processing groups at a frequency of 1 per 25 dosimeters.

## Slide 30
(14) DOE-STD-1095 states that Quality Control and unirradiated dosimeters shall be used routinely to identify reader processing problems. Unirradiated and quality control dosimetry frequency shall be determined based upon the total number of dosimeters processed, equipment stability, type of quality control checks, or other suitable method. The practice used by XYZ (XYZ procedure QA-1-009 step 7.2.7) is that only irradiated Quality Control cards (QC) are interspersed in dosimeter processing groups at a frequency of 1 per 25 dosimeters.4.8.2(d) CONCERN

## Slide 31
(15) Between April 1, 2020, through March 31, 2021, there were four incidents involving at least 13 QC dosimeters and 45 worker dosimeters that were reported as failed analyses due to a reader failure. However, no corrective action request and report (CARR) or an internal nonconformance report were opened to investigate and prevent recurrence. In June 2021, and reader failure incident was finally investigated and corrective actions implemented. Section E in XYZ Quality Assurance Plan (XYZ-QS-B-oo1 rev 36) requires the completion of an internal nonconformance report when laboratory control samples do not meet acceptance criteria specified in the applicable SOP.

## Slide 32
(15) Between April 1, 2020, through March 31, 2021, there were four incidents involving at least 13 QC dosimeters and 45 worker dosimeters that were reported as failed analyses due to a reader failure. However, no corrective action request and report (CARR) or an internal nonconformance report were opened to investigate and prevent recurrence. In June 2021, and reader failure incident was finally investigated and corrective actions implemented. Section E in XYZ Quality Assurance Plan (XYZ-QS-B-oo1 rev 36) requires the completion of an internal nonconformance report when laboratory control samples do not meet acceptance criteria specified in the applicable SOP. 4.7(b) CONCERN

## Slide 33
Whole body dosimeters are worn for the following wear periods: quarterly, semiannual, and annual. Per EGAR 2540, "External Dosimetry Lower Limit of Detection for the Landauer lnlight LOR Model 2, 2T, and U Dosimeters" the lower limit of detection is calculated and documented for a quarterly wear period for the following Landauer OSL dosimeters: lnlight LOR Model 2 and 2T dosimeters, using the algorithm, lnlightMod2N_09 version 00101. The site is working on the semi-annual LLD calculation, and it is near completion. However, an LLD for the annual wear period should be conducted. The lower limit of detection documentation for the extremity program references the U-ring rather than the Saturn ring.

## Slide 34
Whole body dosimeters are worn for the following wear periods: quarterly, semiannual, and annual. Per EGAR 2540, "External Dosimetry Lower Limit of Detection for the Landauer lnlight LOR Model 2, 2T, and U Dosimeters" the lower limit of detection is calculated and documented for a quarterly wear period for the following Landauer OSL dosimeters: lnlight LOR Model 2 and 2T dosimeters, using the algorithm, lnlightMod2N_09 version 00101. The site is working on the semi-annual LLD calculation, and it is near completion. However, an LLD for the annual wear period should be conducted. The lower limit of detection documentation for the extremity program references the U-ring rather than the Saturn ring. 4.8.1(c) Concern

## Slide 35
The formal Management and QA Program review performed in February 2023 does not include the necessary attributes to fully meet DOELAP expectations and requirements.

## Slide 36
The formal Management and QA Program review performed in February 2023 does not include the necessary attributes to fully meet DOELAP expectations and requirements. 4.3(g) Concern

## Slide 37
Senior management was listed as the responsible manager on the management assessment conducted in February 13-17 of 2023. However, despite the QAP stating that XXX will perform the formal assessment in section 12.4, the management assessment was not conducted by the technical lead, QA lead, and member of senior management and does not meet the requirements listed in the standard. The assessors noted that the QAP identifies the specific items of 4.3(g) for the management assessment and that the management assessment conducted in 2021 met the specific items of 4.3(g) for the previous cycle but was not conducted during the current accreditation cycle. 4.3(g)

## Slide 38
Senior management was listed as the responsible manager on the management assessment conducted in February 13-17 of 2023. However, despite the QAP stating that XXX will perform the formal assessment in section 12.4, the management assessment was not conducted by the technical lead, QA lead, and member of senior management and does not meet the requirements listed in the standard. The assessors noted that the QAP identifies the specific items of 4.3(g) for the management assessment and that the management assessment conducted in 2021 met the specific items of 4.3(g) for the previous cycle but was not conducted during the current accreditation cycle. 4.3(g)
Concern, 4.3(g)

## Slide 39
Corrective Action Plan for Concern #1
1. Conduct the mid-cycle management assessment during the next accreditation cycle.
2. Revise EDG-QAP, Quality Assurance Plan, by October 1, 20XX, to:
a. Include primary responsibility and approval for the External Dosimetrist for the Management Review. Primary responsibility and approval for the QA lead and member of senior management is already included.
b. Include the Management Review Agenda Template.

## Slide 40
Although records are maintained in accordance with applicable requirements, the XXX should consider improved storage of records. Numerous hard copy original records are kept in the Dosimetry Processing Lab and long- term storage of XXX records may not adequately protect them from water damage in the event of a fire.

## Slide 41
Although records are maintained in accordance with applicable requirements, the DP should consider improved storage of records. Numerous hard copy original records are kept in the Dosimetry Processing Lab and long- term storage of DP records may not adequately protect them from water damage in the event of a fire. Observation

## Slide 42
The quality assurance manual is incorporated as section 9 of the program technical basis document. The technical basis document has not been updated since 2017. The TBD and QA manual are missing key elements, especially those related to program quality and QA lead responsibilities, as well as the Training program process. 4.2(c) “The QA Lead shall have responsibility and authority to implement the quality assurance program.”

## Slide 43
The quality assurance manual is incorporated as section 9 of the program technical basis document. The technical basis document has not been updated since 2017. The TBD and QA manual are missing key elements, especially those related to program quality and QA lead responsibilities, as well as the Training program process. 4.2(c) “The QA Lead shall have responsibility and authority to implement the quality assurance program.” 4.1(a), 4.2(c), 4.2(d), 4.2(e), 4.2(h), 4.2(i) Deficiency

## Slide 44
Blind audit failures are not being formally documented as an issue adverse to quality and do not include a broader evaluation of program impacts.

## Slide 45
Blind audit failures are not being formally documented as an issue adverse to quality and do not include a broader evaluation of program impacts. 4.8.1(e) Concern

## Slide 46
The site performs regular blind audit testing in the categories to which the program is accredited. Blind audit dosimeters and performance reports are provided by PNNL. Currently the program does not trend the blind testing results or produce formal comparisons to historical performance. The program may benefit from identifying trends in the blind audit testing results, e.g. consistent high or low bias in a particular category or source.

## Slide 47
The site performs regular blind audit testing in the categories to which the program is accredited. Blind audit dosimeters and performance reports are provided by PNNL. Currently the program does not trend the blind testing results or produce formal comparisons to historical performance. The program may benefit from identifying trends in the blind audit testing results, e.g. consistent high or low bias in a particular category or source. Observation